Chapter 1 HL7 Communication Protocol Overview

1.1 Description.

HL7 is an electronic data interchange standard for providing care to hospitalized patients. Originally defined in the United States, it has been adopted by many countries. This interface is defined based on HL7 v2.3.1, please refer to HL7 Interface Standards Version 2.3.1 for details;

1.2 Data transmission protocol hierarchy

The data transfer process is: host application -> client HL7 protocol -> TCP/IP protocol -> LIS system;

The instrument application program collects and processes the electronic data of the samples, and then passes them to the HL7 protocol layer; the HL7 protocol packages the data collected and processed by the host computer and passes them to the TCP/IP protocol layer; the TCP/IP protocol encapsulates the data into byte-stream information and passes them; the LIS system accepts the byte-stream data and then carries out the reverse decoding to get the data of the samples;

1.3 Data transmission protocol flow

Each flow of the instrument and exception handling are as follows:

After the instrument starts up, it enters the connection phase, sending TCP connection application to the LIS system server according to the configuration; if the connection fails or disconnects, it will try to connect again with a connection interval of 1 second; after the connection is successful, it monitors the connection survival status at intervals according to the set heartbeat value to ensure that the data can be sent at any time;

Upload sample detection information: the host computer program experiments to get sample experimental data, after confirmation under the initiative to send a message, send and wait for a confirmation message, if the send error or not receive a confirmation message within 5 seconds, it means that the send failed to trigger the send again, if three consecutive times are unsuccessful, give up sending the report exception;

When the exception occurs, network detection can be performed; the instrument will actively disconnect the current connection and retry to establish a connection with the LIS system server according to the settings;

1.4 Message Encoding Format

The TCP/IP protocol is a byte stream protocol and does not provide message boundaries. Messages are encapsulated using HL7's Minimal Lower Layer Protocol (MLLP). The MLLP protocol is defined by the HL7 standard specification, and the HL7 message encapsulation method adopts a single character at the beginning of the message and a double character at the end of the message.The characters used by the HL7 interface are the default characters of the HL7 standard.

Beginning character: hexadecimal <0B>

End character: hexadecimal <1C><0D>

The message is encoded in UTF-8 format. All sent and received strings are UTF-8 encoded strings;

Chapter 2 HL7 Communication

This protocol is used to transmit sample test results and sample information from the instrument to the LIS system;

2.1 Message Syntax

Each HL7 message consists of a number of message segments, each ending in a character.

Each message segment consists of a three-character segment name and a variable number of fields, each consisting of components and subcomponents.

MSH|^~\&|||||||ORU^R01|RD02|P|2.3.1||||||UTF-8

The MSH message segment defines the separator for fields, components, and subcomponents.

The five characters following the MSH define the separator used to distinguish between fields, components, and subcomponents. Although these characters can be any non-text characters, the HL7 standard recommends the characters in the following table:

Format 1 Character table

| Delimiter | Value | Remarks |
| --- | --- | --- |
| Segment terminator | (hex 0D)Carriage return character | Terminates a segment record; this value cannot be changed depending on the application. |
| Field Separator | ¦ | Used to separate two adjacent data fields in a segment. The field separator ¦ is used to separate two adjacent data fields in a message field. It can also be used to separate the first data field message field ID number in the same message field. |
| Component separator | ^ | Separates adjacent components in a data field where allowed. |
| Subcomponent separator | & | Separates adjacent subcomponents of a data field where allowed. If there are no subcomponents, the character is ignored. |
| Repeat separator | ~ | Separates multiple occurrences of a field where allowed. |
| Escape character | \ | The scape character is used in fields represented by the ST, TXF, or FT data types, or in the data (fourth) component for the ED data type. If the escape character is not used in the message, the character is ignored. However, if the subcomponent is used in the message, the character must be present. |

The first field of the MSH contains the individual separators. Some of the subsequent fields are empty because they are optional and are not used by the HL7 interface.

For any given message, the message segments following the MSH message segment have a fixed order of occurrence, which is described in the following sections, using these syntactic constructs to represent the order of the message segments:

Segments appearing inside [] are optional.

Segments within {} may be repeated one or more times.

The syntax for each message segment is described in the Appendix;

  - 2.2 Unsolicited - ORU^R01 message

The instrument actively sends sample test results, sample information and QC results to the LIS system, as shown in Figure 1.

Instrument terminals

LIS system

ORU^R01

ACK^R01

Figure 1 Schematic diagram of the communication process of test results

Unsolicited transmission of an observation message is sent using the ORU (Unsolicited transmission of an observation message) message;

The following messages are included:

1. patient information needed for network data

2.Bed information

3. parameter and device settings

The basic format is as follows:

ORU ^R01 message:

MSH message header, including communication information such as message number, send time, message separator, and encoding method

{

PID Basic patient information, including patient's name, gender, medical record number, date of birth, etc.

[PV1] Patient visit information, including patient type, department, bed number, fee, etc.

{

OBR Sample information, including sample number, examiner, test time, etc. {[OBX]}

{[OBX]} test data items, including test parameter results and data related to working mode tests, etc.

}

}

ACK^R01 message:

MSH Message Header

MSA Message Acknowledgement describing whether the communication message was successfully received or not

    - 2.2.1 Specific Message Segments

MSH Message Segment

All messages begin with the MSH end.

Example：

MSH|^~\&|GZ-8 Plus|123456789v|||20230131110420||ORU^R01|202301311104200001|0|2.3.1|||AA||CHA|UTF-8

| No. | Field Name | Example | Description Description |
| --- | --- | --- | --- |
| 1 | Field separator | ¦ | Separator between the field ID and the first field, all subsequent field separators continue to use the new separator. |
| 2 | Encoded characters | ^~\& | Declare component separator, repeat separator, escape separator and subcomponent separator. |
| 3 | Sending application | GZ-8 Plus | Sending device name |
| 4 | Sending device | 123456789v | Instrument serial number |
| 5 | Message date and time | 20230131110420 | Generation time |
| 6 | Message type | ORU^R01 |  |
| 7 | Message control ID number | 202301311104200001 | Message Unique Identifier |
| 8 | Sample type | 0 | Sample-0,QC-2 |
| 9 | HL7protocol version | 2.3.1 |  |
| 10 | Message reply type | AL | AL reply required AA no reply required |
| 11 | Country code | CHA | China |
| 12 | Character encoding | UTF-8 |  |

Format 2 MSH Field Definition Table

PID Message Field

The PID message segment is used to identify the patient in an ORU message.

Example：

PID|cyq6||202108301530001|Toxoplasma gondii|Xiaohua||19941009|M

| Serial No. | Field Name | Example | Description |
| --- | --- | --- | --- |
| 1 | ID Setting | cyq6 | Patient number |
| 2 | Sample Number/QC File Number | 20210806001 | Used as the sample number in sample test result messages. |
| 3 | Case number/QC lot number | 202108301530001 | Used as the patient record number in sample test result messages. In QC messages, used to indicate the QC lot number. |
| 4 | Test item | Toxoplasma gondii |  |
| 5 | Patient name | Xiaohua |  |
| 7 | Actual date of birth/QC expiration date | 19941009 | In sample messages, used as date of birth time. In the QC message, used as the QC expiration date |
| 8 | Sex | M | M for male, F for female, O for other |

Table 3 PID field definition table

PV1 Message Segment

PV1 message segment is used in ORU messages to indicate patient bed information.

Example：

PV1|12|I |Internal Medicine||||||||

| No. | Field Name | Example | Description |
| --- | --- | --- | --- |
| 1 | ID Setting | 1 | Patient number |
| 2 | Patient Category | I | B - Maternity, E - Emergency, I - Inpatient, N - Not Applicable, O - Outpatient, R - Repeat; |
| 3 | Patient Location | Internal Medicine | Patient's specific location |

FOrmat 4 PV1Field Definition Table

OBR Message Segment

The OBR message segment is used in ORU messages to indicate important information about the detection experiment;

Example:

OBR|1|||||20210807110000|20210807181945|||Xiaohua||||20170706090000||||||||||||||||||GZYL

| No. | Field Name | Example | Description |
| --- | --- | --- | --- |
| 1 | ID Setting | 1 | Experiment number |
| 2 | Request date time | 20210807110000 | Time of sending inspection |
| 3 | Inspection date time | 20210807181945 | Inspection time |
| 4 | Collector identification | Xiaohua | Sender |
| 5 | Specimen collection time | 20210807090000 | Sampling time |
| 6 | Interpreter of primary results | GZYL |  |

Format 5  OBR Field Definition Table

OBX Message Segment

The OBR message segment is used in ORU messages to represent instrument experiment data;

Example：

OBX|0|STR|DATA| |20210108408^0^6^23^6^^^1^0^20210108^025715^1042^1047^1054^1074^1049^1079^1078^1076^1078^1087^1070^1079^1105^1088^1087^1074^1078^1112^1081^1092^1089^1087^1094^1098^1090^1097^1076^1088^1094^1089^1111^1114^1104^1101^1100^1110^1109^1094^1105^1098^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^-1117.23

| Serial No. | Field Name | Example | Description |
| --- | --- | --- | --- |
| 1 | ID Setting | 0 | Flow number |
| 2 | Value Type | STR | IS-encoded value of user-defined table; ED-encapsulated data; NM-numeric data; ST-string data; TX-text data; |
| 3 | Observation identifier | DATA | Amplification experiment data |
| 4 | Observation value | 20210108408^ … ^-1117.23 | Specific amplified fluorescence data; |

Format 6 OBRField Definition Table

Complete message example:：

MSH|^~\&|GZ-8 Plus|123456789v|||20230131110420||ORU^R01|202301311104200001|0|2.3.1|||AA||CHA|UTF-8

PID|dcv||20220816009|0|dcv||19941009|M

PV1|12| emergency treatment|general medicine||||||||123

OBR|2|||||20230131110420|20230201110420|||abc||||20230131110220||||abc||abc||||NK||||||||GZYL

OBX|0|STR|DATA||20220816009^48^dcv^11^11^^^2^1^20220816^153151^128^128^128^128^128^128^128^128^128^128^128^128^129^130^133^134^135^136^136^138^150^150^153^154^154^164^164^169^169^171^174^174^174^177^178^178^181^181^181^181^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816010^47^dcv^3^3^^^1^1^20220816^153151^114^114^114^114^114^114^114^114^114^114^114^114^118^118^118^118^118^118^122^123^123^127^130^130^130^130^132^133^137^137^137^137^138^142^142^142^142^142^142^142^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816011^42^dcv^10^10^^^2^1^20220816^153151^130^130^130^130^130^130^130^130^130^130^133^133^133^133^133^133^133^133^133^133^134^147^147^147^147^147^152^153^153^153^153^153^162^162^162^164^169^169^169^169^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816012^41^dcv^2^2^^^1^1^20220816^153151^112^112^112^112^112^112^112^112^112^112^112^112^112^112^112^112^112^112^112^115^121^121^122^122^122^129^130^130^130^130^130^130^130^130^134^140^140^140^140^140^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

MSH|^~\&|GZ-8 Plus|123456789v|||20230131110420||ORU^R01|202301311104200002|0|2.3.1|||AA||CHA|UTF-8

PID|dxg||20241028003|0|dxg||19941009|M

PV1|12|emergency treatment|general medicine||||||||123

OBR|2|||||20230131110420|20230201110420|||abc||||20230131110220||||abc||abc||||NK||||||||GZYL

OBX|0|STR|DATA||20241028003^85^dxg^14^14^^^2^1^20220816^153151^124^124^124^124^124^124^124^124^124^124^126^127^131^135^136^138^142^142^142^142^146^150^156^156^158^159^159^159^159^159^159^159^160^163^163^163^163^163^163^163^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816004^84^dxg^6^6^^^1^1^20220816^153151^108^108^108^108^108^108^108^108^108^108^108^112^112^115^115^122^123^123^123^123^123^123^123^123^123^123^125^127^127^127^127^131^136^136^136^136^136^140^141^141^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816005^77^dxg^13^13^^^2^1^20220816^153151^120^120^120^120^120^120^120^120^120^120^120^120^120^120^120^120^123^126^128^128^128^130^134^134^134^134^134^134^134^135^140^143^144^148^148^151^155^155^155^155^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816006^76^dxg^5^5^^^1^1^20220816^153151^106^106^106^106^106^106^106^106^106^106^106^106^106^106^106^106^106^110^114^114^114^114^116^116^116^116^116^116^119^121^121^121^121^121^121^122^122^125^127^127^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816007^88^dxg^12^12^^^2^1^20220816^153151^116^116^116^116^116^116^116^116^116^116^120^120^120^120^120^120^120^120^125^129^129^132^135^135^138^160^160^160^160^160^160^160^160^160^160^160^160^160^160^160^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

OBX|0|STR|DATA||20220816008^87^dxg^4^4^^^1^1^20220816^153151^104^104^104^104^104^104^104^104^104^104^104^105^106^110^110^110^110^116^117^117^117^117^117^117^117^117^117^117^117^117^119^120^121^121^121^121^121^121^128^128^41^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0^0

Where the valid data are MSH data and OBX data;

Where MSH data needs to be parsed for “instrument serial number”

OBX data is all the data of one experiment, up to 16 entries;

Please refer to the field description document for the specific composition and value range of MSH data and OBX data;